## Sheet1
Sub System | Problem | Interlock 1 | Interlock 2 | Interlock 3 | Interlock 4 | Interlock 5 | Error Code 1 | Error Code 2 | Error Code 3 | Resolution 1 | Resolution 2 | Resolution 3 | Resolution 4 | Resolution 5 | Resolution 6 | Resolution 7
Beam | Output too high. | [8 empty] | Ion chamber
Beam | GFIL Interlock | GFIL | [7 empty] | Filament PS | Gun Pulse Driver
Beam | GFIL Interlock | GFIL | [7 empty] | Gun Pulse Driver
Beam | Yield monitor error | YIELD | [7 empty] | Target
Beam | Underdose | UDRS | [7 empty] | Adjusted Gun HV
Beam | GFIL Interlock | GFIL | [7 empty] | Hot Deck PS | Pulser PCB | Filament PS
Beam | GFIL Interlock | GFIL |  |  |  |  | HD AD+ |  |  | LV PS PS2 | Filament PS | Gun Pulse Driver
Beam | Underdose | UDRS1 | UDRS2 | [6 empty] | Cold Deck BP | Gantry Patch Panel | w13 | w15 | w30 | w32 | w88
Beam | GFIL Interlock | GFIL |  |  |  |  | HVPS+ | FIL |  | Filament PS | Gun Pulse Driver
Beam | Yield monitor error | YIELD | [7 empty] | Target
Beam | BGM Fault | BGM |  |  |  |  | 210005 |  |  | Adjusted AFC
Beam | GFIL Interlock | GFIL |  |  |  |  | HD AD+ |  |  | Hot Deck PS
Beam | GFIL Interlock | GFIL | [7 empty] | Gun Pulse Driver | Hot Deck PS
Beam | GFIL Interlock | GFIL | [7 empty] | Emergency Off Reset
Beam | GFIL and HVCB Faults | GFIL | HVCB | [6 empty] | Hot Deck BP
Beam | GFIL Interlock | GFIL | [7 empty] | Adjusted Gun HV
Beam | GFIL Interlock | GFIL | [7 empty] | Emergency Off Reset
Beam | GFIL Interlock | GFIL | [7 empty] | Filament PS | Gun Pulse Driver
Beam | GFIL Interlock | GFIL |  |  |  |  | AD Overcurrent |  |  | Hot Deck BP | LV PS PS2
Beam | GFIL Interlock | GFIL | [7 empty] | Gun Controller
Beam | GFIL Interlock | GFIL | [7 empty] | Gun Pulse Driver | Filament PS
Beam | GFIL Interlock | GFIL | [7 empty] | Filament PS | Gun Pulse Driver | Aztec PS
Beam | GFIL Interlock | GFIL |  |  |  |  | HD AD+ | HV OFF | EE ROM | HVPS
Beam | GFIL Interlock | GFIL |  |  |  |  | HD TEMP | HD AD+ |  | LV PS PS2
Beam | GFIL, UDRS, VAC 2 | GFIL | UDRS | VAC2 | [5 empty] | Tuned Machine
Beam | GFIL Interlock | GFIL | [7 empty] | Gun Pulse Driver | Filament PS
Beam | GFIL Interlock | GFIL |  |  |  |  | HD AD+ |  |  | Reseat Transformer AC Input Cable | Reseat Cold Deck PS
Beam | GFIL Interlock | GFIL | [7 empty] | Gun Pulse Driver | Filament PS
Beam | GFIL Interlock | GFIL |  |  |  |  | HD AD+ |  |  | Filament Harness
Beam | GFIL Interlock | GFIL | [7 empty] | Filament PS
Beam | GFIL Interlock | GFIL | [7 empty] | Reseat Cold Deck PS
MLC | MLC stuck leaf | MLC | [7 empty] | MLC Motor
MLC | MLC stuck leaf | MLC | [7 empty] | MLC Motor | Leaf Nut | Leaf Screw | Leaf Spring
MLC | MLC Initialization Error | MLC | INIT | SFB | [5 empty] | SoftPot | MLC Motor | Leaf Nut | Leaf Screw | Leaf Spring
MLC | MLC Initialization Error | MLC | INIT | [6 empty] | SoftPot
MLC | MLC Initialization Error and Leaf Stuck | MLC | INIT | [6 empty] | MLC Motor
MLC | MLC Initialization Error | MLC | INIT | [6 empty] | MLC Motor
MLC | MLC Initialization Error | MLC | INIT | [6 empty] | MLC Motor | Leaf Nut | Clean MLC Leaf
MLC | MLC losing connection, treatment invalid | MLC |  |  |  |  | Leaf Stall |  |  | MLC Motor
MLC | MLC Leaf Stall | MLC |  |  |  |  | Leaf Stall |  |  | MLC Motor
MLC | MLC stuck leaf | MLC | [7 empty] | MLC Motor
MLC | MLC Leaf Error | MLC | [7 empty] | MLC Motor
MLC | MLC stuck leaf | MLC | [7 empty] | MLC Motor
MLC | MLC stuck leaf | MLC | [7 empty] | SoftPot | Leaf Spring
MLC | MLC Initialization Error | MLC | INIT | [6 empty] | Reset Aux Electronics Power
MLC | MLC Initialization Error | MLC | INIT | [6 empty] | SoftPot
MLC | MLC, COL error | MLC | COL | [6 empty] | MLC Motor | SoftPot
MLC | MLC Leaf Stopping Intermittently | MLC | SFB | [6 empty] | SoftPot
MLC | MLC Leaf Stall | MLC |  |  |  |  | Leaf Stall |  |  | Clean MLC Leaf
MLC | MLC Won't go to plan, Leaf Failure | MLC | [7 empty] | Mosaiq sending wrong MLC plan
MLC | MLC Leaf Error | MLC | [7 empty] | MLC Motor | Leaf Nut | Lead Screw | Clean MLC Leaf
MLC | MLC Stuck Open | MLC | [7 empty] | Reinitialize MLC
MLC | MLC Interlock | MLC | [7 empty] | MLC Motor | Leaf Nut
MLC | MLC Initialization Error | MLC | INIT | [6 empty] | MLC Motor
MLC | MLC stuck leaf | MLC | [7 empty] | SoftPot
MLC | MLC stuck leaf | MLC | [7 empty] | MLC Motor | Clean MLC Leaf
MLC | MLC Secondary Feedback Error | MLC | SFB | [6 empty] | SoftPot | SFB PCB
MLC | MLC Secondary Feedback Error | MLC | SFB | [6 empty] | Clean MLC Leaf
Beam | Underdose in morning | UDRS | [7 empty] | RF Driver
Beam | Underdose in morning | UDRS | UDRS1 | UDRS2 | [5 empty] | Emergency Off Reset
Beam | Underdose in Photons | UDRS | [7 empty] | 3 in 1 Power Supply
Beam | Underdose in 9e | UDRS | [7 empty] | Tuned Machine
Beam | Underdose | UDRS | [7 empty] | Gun Pulse Driver
Beam | Underdose all energies | UDRS | [7 empty] | Tuned Machine
Beam | Underdose | UDRS | [7 empty] | K2 Relay
Beam | Underdose | UDRS | UDRS1 | UDRS2 | [5 empty] | W109
Beam | Underdose in Electrons | UDRS | [7 empty] | Gantry Patch Panel | Gantry Wind Up Bundle
Beam | Underdose when gantry 180 degrees | UDRS | [7 empty] | W32 | Reset Gun Pulse Driver
Beam | Underdose in 15MV in morning | UDRS | [7 empty] | Gun Controller Grid Voltage | Tuned Machine
Beam | Underdose | UDRS | [7 empty] | Cold Deck BP | Gun Crowbar Shockmounts
Beam | Underdose in 6MV and 6e | UDRS | [7 empty] | Cold Deck PS
Beam | Underdose | UDRS | [7 empty] | W30 | Tuned AFC
Beam | Underdose in 8MV | UDRS | [7 empty] | AFC Cables | Tuned AFC
Beam | Underdose | UDRS | [7 empty] | Thyratron
Beam | Undrdose in 6MV | UDRS | [7 empty] | PFN Servo PCB
Beam | UDRS and HWFA faults | UDRS | HWFA | [6 empty] | DQing Thyratron
Beam | Underdose | UDRS | [7 empty] | PWM PCB
Beam | Underdose | UDRS | UDRS1 | UDRS2 | [5 empty] | Gantry Wind Up Bundle
Beam | Underdose when rotating gantry | UDRS | [7 empty] | Gantry Patch Panel
Beam | Underdose | UDRS | EXQ2 | [6 empty] | Ion chamber cable | B20 SI PCB
Beam | Underdose in Electrons | UDRS | [7 empty] | AFC Control PCB
Beam | Underdose in 6MV | UDRS | UDRS1 | UDRS2 | [5 empty] | Thyratron Grid Control PCB | Fault Conditioning PCB
Beam | Underdose on EDW | UDRS | [7 empty] | Beam Steering
Beam | Underdose | UDRS | [7 empty] | PWM PCB
Beam | Underdose | UDRS | [7 empty] | Thyratron Grid Control PCB | Phase Shifter
Beam | UDRS with IMRT and Gantry rotation | UDRS | [7 empty] | W30 | AFC Coaxial Cables
Beam | Underdose | UDRS | UDRS1 | UDRS2 | [5 empty] | Tuned Machine
Beam | Underdose | UDRS | [7 empty] | Cold Deck BP
Beam | Underdose in 6MV when rotating gantry | UDRS | [7 empty] | W88 | W13 | W15 | W57 | Cold Deck BP
Beam | Underdose | UDRS | [7 empty] | W30
Beam | Underdose | UDRS | EXQ2 | [6 empty] | Ion chamber cable
Beam | Underdose | UDRS | [7 empty] | Gun
Beam | Underdose | UDRS | [7 empty] | Gantry Patch Panel | W15
Beam | Underdose | UDRS | [7 empty] | PWM PCB
Beam | Underdose | UDRS | GFIL | [6 empty] | HVPS
Beam | Underdose in Photons | UDRS | [7 empty] | 3 in 1 Power Supply
Beam | Underdose | UDRS | [7 empty] | RF Driver
Beam | Underdose | UDRS | UDRS1 | UDRS2 | [5 empty] | Tuned Machine
Beam | Underdose when gantry 180 degrees | UDRS | UDRS1 | UDRS2 | [5 empty] | Gantry Wind Up Bundle
Beam | Underdose in 6MV | UDRS | [7 empty] | Thyratron
Beam | Underdose | UDRS | UDRS1 | UDRS2 | [5 empty] | AFC Control PCB
Beam | Underdose | UDRS | [7 empty] | Tuned Machine
Beam | Underdose | UDRS | [7 empty] | Beam Hold Interface PCB
Beam | Underdose in Photons | UDRS | [7 empty] | W15 | Gantry Patch Panel
Water | Continuous flow fault | FLOW | [7 empty] | Flow Restrictor | Flow Switch
Water | FLOW interlock | FLOW | [7 empty] | Cleaned Flow Switch
Water | FLOW interlock | FLOW | [7 empty] | STB-WA-515
Water | FLOW interlock | FLOW | [7 empty] | Cleaned Flow Switch
Water | FLOW interlock | FLOW | PUMP | [6 empty] | Cleaned Filter
Water | FLOW interlock | FLOW | PUMP | KFIL | [5 empty] | Chiller
Water | Intermittent FLOW faults | FLOW | [7 empty] | W32
Water | FLOW interlock | FLOW | [7 empty] | W32
Water | Intermittent FLOW faults | FLOW | [7 empty] | Flow Restrictor
Water | FLOW interlock | FLOW | [7 empty] | Flow Transducer
Water | FLOW interlock | FLOW | [7 empty] | Filled Oil
Beam | BMAG interlock when rotating | BMAG | [7 empty] | Repair BMAG cable insulation
Beam | BMAG interlock | BMAG | [7 empty] | Adjusted BMAG voltage
Beam | BMAG interlock | BMAG | [7 empty] | Adjusted BMAG voltage
Beam | BMAG and HVCB interlocks | BMAG | HVCB | [6 empty] | BMAG Power Supply
Beam | BMAG interlock when rotating | BMAG | [7 empty] | W49
Beam | BMAG interlock | BMAG | [7 empty] | BMAG Power Supply
Beam | BMAG interlock | BMAG | [7 empty] | Input IF 2 PCB | CDRIO PCB
Beam | BMAG and UDRS interlocks | BMAG | UDRS | [6 empty] | PFN Servo PCB
Beam | BMAG interlock | BMAG | [7 empty] | Three Way Valve Actuator